# National Imaging Registry (NIR)

Supplier Guidance for DCB 0160 Clinical Safety Assessment

| Document filename: NIR DCB0160 – Supplier Guidance |  |  |  |
| --- | --- | --- | --- |
| Directorate / Programme | NIR Programme Team | Project | National Image Registry |
| Document Reference |  |  |  |
| Director |  | Status | Draft |
| Owner |  | Version | 0.1 |
| Author | Andrew Harrison ETHOS Ltd. Safety Engineer | Version issue date | XX.XX.2025 |
| This is a controlled document; any printed copies of this document are not controlled. |  |  |  |

Document Management

Revision History

| Version | Date | Summary of Changes |
| --- | --- | --- |
| 0.1 | 15.09.2025 | Initial draft document |

Reviewers

This document must be reviewed by the following people:

| Reviewer name | Title / Responsibility | Date | Version |
| --- | --- | --- | --- |
| Andrew Harrison | Principal Safety Engineer, ETHOS Ltd. | 04/01/2026 | 0.1 |
| Kira Pennington | Delivery manager NIR, NHS England |  |  |
| Savan Shah | Clinical Safety Officer, ETHOS Ltd. | 13/01/2026 | 0.1 |

Approved by

This document must be approved by the following people:

| Name | Title | Date | Version |
| --- | --- | --- | --- |

Glossary of Terms

| AFAP | As Far As Possible – The level risk acceptability criteria as per clinical safety standards |
| --- | --- |
| Clinical Risk | Combination of the severity (consequence) and likelihood (probability) of harm to a patient and the likelihood of occurrence of that harm. |
| Clinical Risk Analysis | Systematic use of available information to identify and estimate a clinical risk. |
| Clinical Risk Control | Process in which decisions are made and measures implemented by which clinical risks are reduced to or maintained within, specified levels. |
| Clinical Risk Estimation | Process used to assign values to the severity (consequence) of harm to a patient and the likelihood (probability) of occurrence of that harm. |
| Clinical Risk Evaluation | Process of comparing a clinical risk against given risk criteria to determine the acceptability of the clinical risk. |
| Clinical Risk Management (CRM) | Systematic application of management policies, procedures, and practices to the tasks of analysing, evaluating, and controlling clinical risk. |
| Clinical Risk Management (CRM) Process | A set of interrelated or interacting activities, defined by NHS England to meet the requirements of this standard with the objective of ensuring clinical safety in respect to the development, deployment and use of NIR. |
| Clinical Safety | Freedom from unacceptable clinical risk to patients. |
| Clinical Safety Officer | NHS Digital accredited clinician (contracting to NHS England), responsible for ensuring the safety of NIR through the application of clinical risk management as set-out in the NHS Digital DCB 0129 Standard requirements. |
| Clinical Safety Case Report (CSCR) | A report that presents the arguments and supporting evidence that provides a compelling, comprehensible, and valid case that a solution is safe for a given application in each environment. |
| ETHOS Ltd. | Clinical Risk Management Health IT subject Matter Experts (Clinical Safety Engineers and Clinical Safety Officers) contracting to NHS England providing Health IT Clinical Safety assurance in respect of NIR. |
| Harm | Death, physical injury, psychological trauma and/or damage to the health or well-being of a patient. |
| Hazard | Potential source of harm to a patient. |
| Hazard Log | A mechanism for recording and communicating the on-going identification of System hazards associated with a Health IT System. |
| Health IT System | Product (such as NIR) used to provide electronic information for health or social care purposes. The product may be hardware, software, or a combination. |
| Residual Clinical Risk | The clinical risk derived during clinical risk estimation for the design. |
| Intended Use | Use of System in accordance with the specifications, instructions and information provided by the Manufacturer to its clients for its intended use. |
| Likelihood (probability) | Measure of the occurrence of harm. |
| Manufacturer | Health Navigator, with responsibility for the design and support of the NIR solution. |
| Patient Safety | Freedom from harm to the patient. |
| Residual Clinical Risk | Clinical risk remaining after the application of risk control measures. |
| Severity (Consequence) | Measure of the possible consequences of a hazard. |

Contents

National Imaging Registry (NIR)	1

National Imaging Registry (NIR) Supplier Guidance for DCB 0160 Clinical Safety Assessment	6

1.	Purpose of this Guidance	6

2.	Why a DCB 0160 Safety Assessment is Required	6

3.	Expected Clinical Safety Governance – Systems and Roles	6

4.	Expected Clinical Safety Governance Arrangements for NIR Integrations are as follows:	6

5.	Source of Hazards and Relationship to DCB 0129	8

6.	Using the DCB 0160 CSCR Template	9

7.	Using the Supplier Hazard Log	9

8.	Conducting the Safety Assessment	9

9.	Submission and Assurance	10

Appendices	11

Appendix A – Risk Matrices and Definitions	11

Appendix B – NIR Deployment Controls and Safety Recommendations (DCB 0160)	13

Appendix C – Detailed Data Flow	15

Appendix D – Standards Compliance	16

# National Imaging Registry (NIR) Supplier Guidance for DCB 0160 Clinical Safety Assessment

This guidance document is provided to support suppliers integrating systems with the National Imaging Registry (NIR) in completing a DCB 0160 Clinical Safety Assessment. It should be read alongside the NIR DCB 0160 Clinical Safety Case Report (CSCR) template, the exemplar supplier CSCR, and the NIR DCB 0129 Clinical Safety Case.

## Purpose of this Guidance

This document explains why a supplier clinical safety assessment is required, how it fits within both the NIR programme and NHS statutory requirements, and how suppliers are expected to use the provided DCB 0160 templates to produce a compliant safety submission.

## Why a DCB 0160 Safety Assessment is Required

DCB 0160 is a mandatory NHS clinical risk management standard applicable to the deployment and use of health IT systems. While the National Imaging Registry (NIR) platform has been assured under DCB 0129 for design and manufacture, suppliers integrating with NIR introduce additional risks through configuration, interoperability, training, and operational use. A DCB 0160 assessment is therefore required to ensure these deployment-related risks are identified and controlled.

From a project perspective, such an assessment ensures consistent onboarding of suppliers and provides assurance that integration does not compromise patient safety. From an NHS governance perspective, compliance with DCB 0160 is a statutory requirement under the Health and Social Care Act 2012.

## Expected Clinical Safety Governance – Systems and Roles

In line with DCB 0160, Suppliers are expected to operate an appropriate Clinical Safety Management System (CSMS) and to appoint a suitably qualified Clinical Safety Officer (CSO) . The CSO is responsible for oversight of the safety assessment and for approving residual risk acceptability.

In addition, suppliers should involve subject matter experts as appropriate, including:

- Clinical representatives familiar with imaging workflows

- Technical architects responsible for NIR integration

- Information governance and security specialists

Evidence of governance arrangements, roles, and responsibilities should be referenced within the DCB 0160 CSCR.

## Expected Clinical Safety Governance Arrangements for NIR Integrations are as follows:

Baseline safety ownership (supplier systems)

Suppliers of clinical systems already deployed within healthcare organisations (e.g. PACS, RIS, VNA, EPR systems) retain responsibility for the clinical safety of their own products under DCB 0129.

Each supplier maintains:

- A DCB 0129 Clinical Safety Case for their system

- A hazard log covering risks inherent to the design, configuration, and intended use of their product

This responsibility does not transfer when those systems are integrated with NIR.

In other words integration with NIR does not invalidate or replace existing supplier DCB 0129 obligations.

NIR system design safety (central responsibility)

NIR, as a nationally designed system, has its own DCB 0129 Clinical Safety Case that covers:

- The design, architecture, and intended use of NIR

- Hazards arising from:

  - Central data aggregation

  - Indexing, querying, and image metadata handling

  - National-level workflows and access patterns

Assumptions about how NIR will be:

- Integrated

- Used

- Governed locally

The NIR DCB 0129 does not:

Re-assess the internal safety of PACS/RIS systems

Take ownership of deployment-specific or local integration risks

Local integration and deployment safety (DCB 0160)

Each healthcare organisation integrating local systems with NIR is required to assure safe deployment and use through DCB 0160.

Where suppliers undertake integration activities on behalf of healthcare organisations, they should do so within a clearly defined governance framework.

Key points:

- DCB 0160 applies to the integration to NIR (deployment), not the underlying system designs

- The DCB 0160 safety case must address:

  - Data flows between local systems and NIR

  - Configuration choices

  - Workflow changes

  - User access and operational context

The DCB 0160 safety case remains under local governance of the healthcare organisation, even if authored by a supplier.

Supplier role in DCB 0160 (on behalf of healthcare organisations)

Where suppliers carry out integration work, they may:

- Author or contribute to the DCB 0160 documentation

- Identify and assess integration-specific hazards

- Propose mitigations and controls

However, the healthcare organisation retains accountability for accepting residual risk and approving the DCB 0160 safety case.

This ensures:

- Clinical risk acceptance remains with the organisation delivering care

- Suppliers do not self-accept deployment risk on behalf of the NHS

Hazard ownership and boundaries, i.e. the governance model relies on clear hazard boundaries:

| Hazard type | Owned under |
| --- | --- |
| Inherent risks of product e.g. PACS / RIS design. | Supplier DCB 0129 |
| Inherent risks of NIR design | NIR DCB 0129 |
| Risks introduced by integration between systems | DCB 0160 (local) |
| Risks due to local configuration and workflows | DCB 0160 (local) |
| Changes to supplier system behaviour | Supplier DCB 0129 (updated if required) |

The principle being that no single safety case is expected to cover all risks end-to-end. Instead, safety is assured through linked, complementary safety cases.

## Source of Hazards and Relationship to DCB 0129

The hazards addressed in the supplier DCB 0160 assessment can be derived from the NIR DCB 0129 Clinical Safety Case. These hazards reflect known safety risks associated with national imaging discovery and retrieval, such as incorrect patient identification, missing images, delayed access, and unauthorised access.

Suppliers are not expected to re-identify or reassess NIR design hazards. Instead, suppliers should focus on how these hazard themes may manifest during deployment and use of their own systems and what additional controls are required.

## Using the DCB 0160 CSCR Template

The DCB 0160 CSCR template has been structured to provide a guide for suppliers for a  deployment safety argument.

If the template is used, Suppliers should:

- Complete all sections of the CSCR in the context of their specific product and deployment model

- Retain the overall structure and headings to ensure consistency across suppliers

- Populate all explicit placeholders with supplier-specific information and evidence

The CSCR should clearly explain how hazards are mitigated through architecture, controls, testing, deployment, and ongoing monitoring.

## Using the Supplier Hazard Log

It is recommended that a DCB 0160 compliant Hazard Log is used. Guidance on the format and content of the Hazard Log  can be found at:

DCB0160: Clinical Risk Management: its Application in the Deployment and Use of Health IT Systems

Also there is a DCB 0129 Hazard Log available for the NIR Design safety assessment which contains guide hazards to consider.

The Hazard Log should be used to document deployment and operational hazards relevant to NIR integration. Hazards should be mapped to the NIR hazard themes provided, where possible, and should reference corresponding DCB 0129 hazards, controls and evidence.

Suppliers should avoid duplicating DCB 0129 design hazards and instead demonstrate how deployment risks are managed within their own systems and processes.

## Conducting the Safety Assessment

Suppliers should conduct the DCB 0160 safety assessment as a structured activity involving both clinical and technical stakeholders. Recommended activities include:

- Review of the NIR DCB 0129 safety case and appendices

- Architectural and data flow review for the supplier system

- Identification of deployment-specific hazards

- Definition and implementation of controls

- Hazard-based testing and validation

The assessment should be proportionate to the complexity and risk of the integration.

Review and Update of the Safety Assessment

The DCB 0160 safety assessment is not a one-off activity. Suppliers are expected to review and update their CSCR and Hazard Log:

- Prior to initial go-live

- Following significant system or configuration changes

- After safety incidents or near misses

- At agreed periodic intervals (recommended at least annually)

Evidence of review and CSO approval should be retained.

## Submission and Assurance

The completed DCB 0160 CSCR and supporting evidence should be submitted as part of the NIR supplier onboarding process. NHS England may review submissions to confirm compliance with DCB 0160 and alignment with NIR safety requirements.

Suppliers should ensure submissions are complete, internally approved, and consistent with other onboarding documentation.

The NIR Clinical Safety Guidance Document is sufficient as a supplier guidance artefact and, when used alongside the NIR DCB0160 CSCR template and a supplier-maintained Hazard Log, provides an appropriate and proportionate framework to support DCB0160-compliant deployment clinical safety assurance.

# Appendices

## Appendix A – Risk Matrices and Definitions

Clinical Risk Management Risk Matrix

| Likelihood | Very High | 3 | 4 | 4 | 5 | 5 |
| --- | --- | --- | --- | --- | --- | --- |
|  | High | 2 | 3 | 3 | 4 | 5 |
|  | Medium | 2 | 2 | 3 | 3 | 4 |
|  | Low | 1 | 2 | 2 | 3 | 4 |
|  | Very Low | 1 | 1 | 2 | 2 | 3 |
|  |  | Minor | Significant | Considerable | Major | Catastrophic |
|  |  | Consequence |  |  |  |  |

Risk Matrix key - Severity

| 5 | Unacceptable level of risk |
| --- | --- |
| 4 | Mandatory elimination or control to reduce risk to an acceptable level |
| 3 | Undesirable level of risk Attempts should be made to eliminate or control to reduce risk to an acceptable level. Shall only be acceptable when further risk reduction is impractical. |
| 2 | Acceptable where cost of further reduction outweighs benefits gained. |
| 1 | Acceptable, no further action required |

Hazard likelihood definitions

| Likelihood Category | Interpretation |
| --- | --- |
| Very high | Certain or almost certain; highly likely to occur |
| High | Not certain but very possible; reasonably expected to occur in the majority of cases |
| Medium | Possible |
| Low | Could occur but in the great majority of occasions will not |
| Very low | Negligible or nearly negligible possibility of occurring |

Hazard Consequence definitions

| Consequence Classification | Interpretation | Number of Patients Affected |
| --- | --- | --- |
| Catastrophic | Death | Multiple |
|  | Permanent life-changing incapacity and any condition for which the prognosis is death or permanent life-changing incapacity; severe injury or severe incapacity from which recovery is not expected in the short term | Multiple |
| Major | Death | Single |
|  | Permanent life-changing incapacity and any condition for which the prognosis is death or permanent life-changing incapacity; severe injury or severe incapacity from which recovery is not expected in the short term | Single |
|  | Severe injury or severe incapacity from which recovery is expected in the short term | Multiple |
|  | Severe psychological trauma | Multiple |
| Considerable | Severe injury or severe incapacity from which recovery is expected in the short term | Single |
|  | Severe psychological trauma | Single |
|  | Minor injury or injuries from which recovery is not expected in the short term. | Multiple |
|  | Significant psychological trauma. | Multiple |
| Significant | Minor injury or injuries from which recovery is not expected in the short term. | Single |
|  | Significant psychological trauma | Single |
|  | Minor injury from which recovery is expected in the short term | Multiple |
|  | Minor psychological upset; inconvenience | Multiple |
| Minor | Minor injury from which recovery is expected in the short term; minor psychological upset; inconvenience; any negligible severity | Single |

## Appendix B – NIR Deployment Controls and Safety Recommendations (DCB 0160)

The Deployment Controls and Safety Recommendations outlined in this DCB 0160 Clinical Safety Case Report for the NIR solution represent risk controls and safety considerations identified during the clinical risk assessment.

| Hazard Group | Individual Hazard IDs | Hazard Sub-Group | Operational Requirements to be Maintained and Recommendations |
| --- | --- | --- | --- |
| Availability & Retrieval Failures | HAZ 001, HAZ 002 | Network & Infrastructure Availability | Organisations must maintain high network uptime and monitor PACS/VNA gateway performance continuously. Network redundancy, load balancing, and failover configurations must be validated and documented. Automated alerts should notify support teams of connectivity issues, and escalation procedures should be triggered for sustained outages. |
| Availability & Retrieval Failures | HAZ 001, HAZ 003, HAZ 004 | Image Retrieval Workflow | System behaviour must ensure complete retrieval of images and related metadata across federated networks. Partial query results must generate explicit user notifications. Retry logic must be implemented, and timeouts configured to prevent misleading outputs. Testing should include simulated retrieval failures. |
| Availability & Retrieval Failures | HAZ 001, HAZ 004, HAZ 002 | Operational Response & Escalation | Local SOPs must define fallback options when image retrieval fails. Clinical and technical users must be trained on escalation routes and how to differentiate between local and remote issues. Tabletop exercises should verify readiness and escalation effectiveness. |
| Data Quality & Patient Identification Errors | HAZ 003, HAZ 005 | Metadata Conformance & Validation | All connected systems must perform schema validation before publishing data to the NIR. Incomplete or malformed metadata must be rejected automatically. Supplier conformance tests during onboarding must validate compliance with IHE XDS and XCA profiles. |
| Data Quality & Patient Identification Errors | HAZ 005 | Patient Identity Management | Patient identifiers must be validated against authoritative sources, and merge/unmerge processes logged. Systems must prevent accidental overwrites or duplicate creation. Regular audits of patient demographic consistency across federated systems are required. |
| Data Quality & Patient Identification Errors | HAZ 003, HAZ 005 | Data Provenance & Traceability | The NIR must maintain the 'single source of truth' principle. Audit logs must document the origin of each dataset and timestamp of registration. Provenance records should be reviewed periodically to ensure traceability and prevent data drift. |
| Security, Privacy & Access Control Failures | HAZ 006 | Identity Federation & Authentication | All NIR users must authenticate via NHS Identity or federated access controls. Access tokens must be time-bound and traceable. Failed authentication attempts should be logged and reviewed for anomalies. |
| Security, Privacy & Access Control Failures | HAZ 006 | Role-Based Access & Least Privilege | Access permissions must be role-based and reviewed quarterly to ensure alignment with job functions. Clinical data should only be accessible to authorised staff, and audit trails must capture all image access events. |
| Security, Privacy & Access Control Failures | HAZ 006 | Encryption & Data Transmission Security | All communication between NIR components and connected systems must use TLS 1.2+ encryption. Certificate management must follow NHS Cyber Security guidelines, with regular rotation and validation. Breach response procedures must be rehearsed annually. |
| Governance & Operational Risks | HAZ 007, HAZ 008 | Governance & Accountability Structures | Each trust must define governance roles for NIR-related incidents and maintain documented accountability structures. Governance meetings should include periodic review of clinical safety reports and escalation trends. |
| Governance & Operational Risks | HAZ 002, HAZ 006 | Incident Escalation & Safety Oversight | Incident management must align with NHS England’s national escalation framework. Clinical Safety Officers must be informed of all incidents with potential patient impact. Incident reporting metrics should be reviewed monthly. |
| Governance & Operational Risks | HAZ 001, HAZ 003, HAZ 004, HAZ 005 | Training & Readiness Testing | Clinical and technical teams must undergo induction training on NIR use, safety principles, and failure procedures. Tabletop and live exercises should validate understanding and readiness at least annually. |
| Interoperability & Protocol Translation Risks | HAZ 003, HAZ 004, HAZ 009 | IHE Profile Conformance (XDS, XCA) | Suppliers must prove conformance with IHE XDS and XCA standards through onboarding testing. Non-conformant implementations must be risk assessed and formally accepted by NHS England. |
| Interoperability & Protocol Translation Risks | HAZ 009 | API & Future Standards Alignment | Future development phases must ensure APIs align with evolving interoperability frameworks. Backward compatibility testing must confirm safe transition between standards. |
| Interoperability & Protocol Translation Risks | HAZ 003, HAZ 004, HAZ 010 | Cross-Vendor Translation & Gateway Behaviour | Gateway components must handle message translation consistently across vendor systems. Monitoring must detect translation errors and latency. Vendors must participate in interoperability testing during each major release. |

## Appendix C – Detailed Data Flow

## Appendix D – Standards Compliance

Referenced Standards and Frameworks within NIR Documentation

| Category | Standard / Framework | Description / Use within NIR |
| --- | --- | --- |
| Document Exchange | IHE XDS.b (Cross-Enterprise Document Sharing) | Core profile for registering and discovering clinical documents. |
| Cross-Community Access | IHE XCA (ITI-38 / ITI-39) | Enables discovery and retrieval of imaging and report data across organisations. Referenced in NIR Error Handling-v33 and Trust Model-v24. |
| FHIR Document Sharing | IHE MHDS (Mobile Health Document Sharing) | FHIR-based standard for federated document access. Supported in parallel with XCA (NIR Product Brief). |
| Imaging Retrieval | DICOM WADO-RS / WADO-WS | Imaging retrieval protocols for RESTful and SOAP-based access to diagnostic images. Mentioned in Trust Model-v24. |
| Security & Identity | WS-Security (OASIS Standard) | SOAP message signing and encryption (for XCA / XDS exchanges). |
|  | SAML 2.0 / XUA (Cross-Enterprise User Assertion) | Provides user identity and authorisation in SOAP exchanges. |
|  | OAuth 2.0 / IUA (Internet User Assertion) | RESTful (FHIR/MHDS) authentication and authorisation mechanism. |
|  | Mutual TLS (mTLS) via NHS Private CA | Enforced at transport layer for system authentication (Trust Model-v24, Security and Compliance Strategy-v5). |
| Data Representation | HL7® FHIR® R4 | Used for MHDS document queries and retrievals. |
| Metadata & Governance | NHS England DDC Interoperability Standards Catalogue | NHS reference for national-level interoperability alignment. |
| Error Handling & Compliance | IHE IT Infrastructure Technical Framework Vol. 3 (ITI TF-3) | Referenced for error response patterns in NIR Error Handling Framework-v33. |

Related Documents

These documents provide additional information and are specifically referenced within this document. Links to documentation have been provided where available and were accurate at the time of issue of this report.

| Ref | Title (NIR Confluence Link) | Version |
| --- | --- | --- |
| 1 | NHS Standard: DCB 0129 – Clinical Risk Management: its Application in the Manufacture of Health IT Systems – Specification | V3.2 |
| 2 | NHS Standard: DCB 0160 – Clinical Risk Management: its Application in the Deployment and Use of Health IT Systems – Specification | V4.2 |
| 3 | DCB 0129 Hazard Log | D0.1 |
| 4 | MHRA Software Classification Guidance: Medical Device Stand-alone Software Including Apps (including IVDMDs), MEDDEV 2.1/6 | 2.1/6 |
| 5 | NHS England Developer Portal (2025): Digital Onboarding Process – https://digital.nhs.uk/developer/guides-and-documentation/digital-onboarding | - |
| 6 | NIR Solution Design Specification (Derived from NIR Product Brief v3 (28 Jun 2025) | v1.5 |
| 7 | NIR Interoperability and Standards Compliance https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1128204655/NIR+Interoperability | Appx E |
| 8 | NIR Observability and Logging Strategy https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1121401300/NIR+Observability+and+Logging+Strategy | Version 14 (12 September 2025) |
| 9 | NIR Error Handling Framework: https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1161766556/NIR+Error+Handling | Version 33 (12 September 2025) |
| 10 | NIR Supplier Onboarding Pack and Supplier Conformance Criteria (SCC) | DOSM |
| 11 | NIR Quality Assurance and Test Strategy https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1098875238/UAT-+Test+Pack | v3.0 |
| 12 | NIR Programme Communications Plan - Governed by NHS England Diagnostics Programme Communications and Engagement Strategy (DDC, 2025) see also - https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1111209985/NIR+-+References | NHS DDC Programme Document, 2025 |
| 13 | Clinical Safety Management System (CSMS) – NHS England Clinical Safety Team - published NHS document governing all DCB 0129/0160 processes and responsibilities. https://www.england.nhs.uk/long-read/digital-clinical-safety-assurance/ | Version 1.2, 03 March 2025 |
| 14 | NIR Clinical Risk Management Plan (CRMP) | v2.0 |
| 15 | NIR Operational Support Model | v2.0 |
| 16 | Incident Management Procedure (IMP) https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1132940087/NIR+Incident+Management+Priority+Level | - |
| 17 | NIR Change Management Process https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1154670369/NIR+-+New+Feature+Requirements+Process also: - https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1187350949/NIR+Key+Operating+Procedures | - |
| 18 | NIR Service Review Reports - Derived from NHS England Service Review and KPI Reporting Framework (DDC, 2025) and NIR Reporting Dashboard v164. | Aligned to NHS DDC Reporting Framework (2025) |
| 19 | NIR Service Desk Process Flow - Based on NHS England DDC Service Operations Handbook (2025) and ITIL v4 Incident and Request Fulfilment workflows. | Aligned to NHS DDC Service Desk Framework (2025) |
| 20 | NIR Monitoring and Alerting Platform - NIR Observability and Logging Strategy v14 (12 Sep 2025); built on AWS CloudWatch, Splunk, and NHS CSOC integration standards (2024). | Derived from NIR Observability v14 (12 Sep 2025) |
| 21 | NIR Interoperability Testing Procedure https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1098875238/UAT-+Test+Pack | - |
| 22 | NIR Confluence Workspace | NIR |
| 23 | NIR Governance Model – Derived from https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1161775785/API+Catalogue+Website+Requirements | - |
| 24 | NIR Clinical Risk Management File (CRMF) https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1195113158/NIR+-+Clinical+Safety+DCB+0129+and+DCB+0160 | - |
| 25 | Clinical Safety Workshop Records https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1208353642/NIR+-+Hazard+Assessment+Workshop+1+30+09+2025 also: https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1208353679/NIR+-+Hazard+Assessment+Workshop+2+10+10+2025 | - |
| 26 | Change and Release Safety Assessments – Derived from: https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1195113171/Governance+R+R+Safety+Plan | - |
| 27 | Medical Device Classification Assessment Form | TBA See Section Medical Device Regulations |
| 28 | NIR Data Protection Impact Assessment (DPIA) | Version 0.5 (8 August 2025) |
| 29 | NIR Reporting and Governance Dashboard https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1149081087/NIR+Reporting+Management+Dashboard%C2%A0-+Requirements+Specification | - |
| 30 | NIR Security Architecture and Controls Framework https://nhsd-confluence.digital.nhs.uk/spaces/CLP/pages/1121401325/NIR+Security+and+Compliance+Strategy | - |

END OF SECTION